## Sheet1
Title | Journal / Conference | Publication date | Summary | Algorithm input | Network input | Secondary information extracted from input (e.g., optical flow) | Network architecture | Network output | Loss function (fidelity term) | Loss function (regularisation term) | Dataset (anatomy, protocol, size and length) | Ground truth acquisition | If training and test data are from the same dataset | Evaluation metric(s) | Expected values and std of metric(s)
Determination of scan-plane motion using speckle decorrelation: Theoretical considerations and initial test | International Journal of Imaging Systems and Technology | December 07, 1998 | Chen et al. proposed estimating scan-plane motion in ultrasound by analysing the rate of change in the correlation of echo signal intensities across B-mode images. | Ultrasound frames | NA | NA | NA | NA | NA | NA | Tissue-mimicking phantom (RMI Model 415,Middleton, WI) | Mechanical scanning with a uniform known 0.2 mm image slice separation. | NA | Image slice separation (step size ) | 0.20±0.01 mm
Sensorless freehand 3-D ultrasound using regression of the echo intensity | Ultrasound in Medicine & Biology | March, 2003 | Prager et al. proposed a sensor-free method for estimating probe trajectory in freehand 3D ultrasound using linear regression of echo-envelope intensity signals, derived from a probabilistic speckle analysis, where the regression gradient directly relates to probe motion. | B-scan images | NA | NA | NA | NA | NA | NA | 1) Tissue-mimicking phantoms: two sets of data, including 50 slices and 100 slices, respectively. 2) A phantom containing spheres (2 mm in diameter); 3) A freehand 3-D US scan of a 16-week-old foetus in utero. | For dataset 1): A rigid mechanical to enable the acquisition of slices with accurate known spacing; For dataset 2): Polaris optical tracker (Northern Digital Inc., Waterloo, Ontario, Canada); For dataset 3): A Fastrak (Polhemus Inc., Colchester, Vermont, Canada) magnetic sensor. | NA | For dataset 1): the relationship between separation of B-scan slices (horizontal axis) and the estimated separation; For dataset 2): consistency of the reconstruction of the sphere locations; For dataset 3): volume of the fetus | For dataset 1): exact number is not provided; For dataset 2): of the sphere centres, 95% are located to within ± 1.5 mm of their correct position in the best-fit grid; For dataset 3): longitudinal direction (about 14% compression); volume error of the fetus: 10 ml.
3-D US frame positioning using speckle decorrelation and image registration | Ultrasound in Medicine & Biology | June, 2003 | Chang et al. presented a 3D ultrasound positioning system that replaced external hardware by combining speckle decorrelation, to estimate relative distances between neighboring frames, with image registration, to refine frame positions against a perpendicular reference image. | Ultrasound frames in two directions (one direction is for reconstruction and one frame along the reference direction for image registration) | NA | NA | NA | NA | NA | NA | Dataset 1): Ball; 55 frames; Dataset 2): Tumor; 42 frames; Dataset 3): Tumor; 32 frames; Dataset 4): Tumor; 116 frames; Dataset 5): Ball; 59 frames. | Steadily moving a probe in a predefined manner, transversely and longitudinally. | NA | 1) The accuracy rates of the ball width and area; 2) Processing time | 1) The accuracy rates of the ball width and area: Dataset 1: for ROI height of 55 and shifting pixel of 1: 98.28%; 99.66%; Dataset 2: for ROI height of 40 and shifting pixel of 1: 94.12%, 89.07%; Dataset 3: for ROI height of 45 and shifting pixel of 1: 97.56%; 95.32%; Dataset 4: for ROI height of 135 and shifting pixel of 1: 97.52%, 92.29%; Dataset 5: 96.55%, 95.27%. 2) Processing time: Dataset 1: 24s; Dataset 2: 16s; Dataset 3: 12s; Dataset 4: 67s; Dataset 5: 24s;
Sensorless freehand 3D ultrasound in real tissue: Speckle decorrelation without fully developed speckle | Medical Image Analysis | April, 2006 | Gee et al. proposed a heuristic technique to quantify the amount of coherency at each point in the B-scans, enabling an adapted elevational decorrelation scheme. | Ultrasound frames | NA | NA | NA | NA | NA | NA | 1) A speckle phantom: elevational scan spacing was 0.04 mm; 100 parallel B-scans; 2) A beef joint: one was acquired parallel to the muscle fibres, the other perpendicular to them; both had an elevational scan spacing of 0.04 mm; 100 parallel B-scans; 3) A stack of turkey fillets lying in a water bath; an elevational scan spacing of 0.02 mm; 100 parallel B-scans. | Displacement was measured using a dial gauge | NA | 1) Distance and angle estimation; 2) Accumulated errors in full 3D reconstructions. | When frame separation is 0.24 mm, tilt = 0, yaw = 0: 1) Phantom dataset: Distance: 0.240±0.011 mm; Tilt: -0.008±0.022 deg; Yaw: -0.000±0.013 deg; 2) Beef dataset 1: Distance: 0.235±0.014 mm; Tilt: 0.016±0.101 deg; Yaw: 0.022±0.027 deg; 3) Beef dataset 2: Distance: 0.244±0.012 mm; Tilt: 0.007±0.157 deg; Yaw: -0.079±0.042 deg; 4) Turkey dataset: Distance: 0.248±0.009 mm; Tilt: -0.056±0.032 deg; Yaw: -0.057±0.026 deg. Actual width for the first three datasets: 3.96 mm, for turkey data set: 1.98 mm; tilt = 0, yaw = 0: 1) Phantom dataset: Width: 3.95 mm; Tilt: -0.06 deg; Yaw: 0.05 deg; 2) Beef dataset 1: Width: 3.84 mm; Tilt: 1.13 deg; Yaw: 0.54 deg; 3) Beef dataset 2: Width: 3.98 mm; Tilt: 0.55 deg; Yaw: -0.85 deg; 4) Turkey dataset: Width: 2.04 mm; Tilt: 0.10 deg; Yaw: -0.07 deg.
Sensorless reconstruction of freehand 3D ultrasound data | Medical Image Computing and Computer-Assisted Intervention | 2006 | Housden et al. presented reconstruction methods that handled irregular frame spacing, non-monotonic out-of-plane motion, and significant in-plane motion. | A sequence of frames | NA | NA | NA | NA | NA | NA | 10 scans from beef joints (5 monotonic scans and 5 non-monotonic scans) | Polaris optical tracking system | NA | Accumulated centre offset | Line plots; exact numbers are not provided.
Subsample interpolation strategies for sensorless freehand 3D ultrasound | Ultrasound in Medicine & Biology | December, 2006 | Housden et al. introduced a novel interpolation strategy for freehand 3D ultrasound and demonstrated its robustness to simultaneous lateral and elevational probe motion. | Ultrasound frames | NA | NA | NA | NA | NA | NA | Speckle phantom: 101 B-scans with accumulated lateral motion being 10 mm; 101 B-scans with the total lateral and elevational motion being 10.6 mm; | Measured using a dial gauge | NA | Accumulated lateral offset and roll | Accumulated lateral length and roll for purely lateral probe motion, Stride 3; Correct: Length: 9.90 mm; Roll: 0.00 deg; Offset Gaussian: Length: 9.80 mm; Roll: -0.31 deg. Accumulated lateral length and roll for combined lateral and elevational probe motion, Stride 3; Correct: Length: 10.50 mm; Roll: 0.00 deg; Offset Gaussian: Length: 10.33 mm; Roll: 0.02 deg. Accumulated elevational length and tilt for combined lateral and elevational probe motion, Stride 3; Correct: Length: 10.50 mm; Tilt: 0.00 deg; Offset Gaussian: Length: 9.90 mm; Tilt: 0.00 deg.
Sensorless reconstruction of unconstrained freehand 3D ultrasound Data | Ultrasound in Medicine & Biology | March, 2007 | Housden et al. presented reconstruction methods for unconstrained freehand 3D ultrasound that addressed challenges such as irregular spacing, intersecting frames, non-monotonic out-of-plane motion, and significant in-plane motion. | A sequence of frames | NA | NA | NA | NA | NA | NA | 1) Speckle phantom: combined nonmonotonic elevational translation (20 mm forward, 10 mm reverse), 10° tilt and 15° yaw. 2) A joint of beef: five linear scans with no intentional change of direction; five linear scans with the direction of motion intentionally reversed approximately half way through the scan; three rotational scans. | Northern Digital Polaris optical tracking system | NA | 1) Cumulative elevational offset at the centre of the frames; 2) Axial rotation reconstructions in terms of the cumulative yaw angle; 3) Cumulative elevational offset for the phantom scan. | Line plots; exact numbers are not provided.
Rotational motion in sensorless freehand three-dimensional ultrasound | Ultrasonics | September, 2008 | Housden et al. analysed the effect of probe rotation on the decorrelation curve for image-based motion tracking and proposed a method to correct these curves using measurements from the orientation sensor for improved reconstruction. | 1) A sequence of frames; 2) The actual tilt angle between the image frames, measured using the sensor | NA | NA | NA | NA | NA | NA | 1) Simulated data 2) Phantom data: four sequences of 15 frames | Phantom data: recorded with known elevational separation and tilt. | NA | The accumulated elevational offset at the frame centre, and the accumulated tilt angle, for each of the reconstructed sequences | 1) Simulated data: Length: Actual: 2.80 mm; Proposed: 2.57 mm; Tilt: Actual: 4.20 deg; Proposed: 3.99 deg; 2) Dataset 1: Length: Actual: 2.80 mm; Proposed: 4.15 mm; Tilt: Actual: 4.19 deg; Proposed: 3.08 deg; 3) Dataset 2: Length: Actual: 2.80 mm; Proposed: 2.93 mm; Tilt: Actual: 4.19 deg; Proposed: 1.96 deg; 4) Dataset 3: Length: Actual: 2.79 mm; Proposed: 3.15 mm; Tilt: Actual: 6.28 deg; Proposed: 4.92 deg; 5) Dataset 4: Length: Actual: 2.79 mm; Proposed: 2.85 mm; Tilt: Actual: 6.28 deg; Proposed: 3.84 deg.
On feature motion decorrelation in ultrasound speckle tracking | IEEE Transactions on Medical Imaging | November 27, 2012 | Liang et al. analysed two types of speckle pattern variations, the geometric transformation and the intensity change of speckle patterns, and demonstrated that a coupled filtering method or an affine warping method can compensate for these effects to provide accurate strain estimation under large tissue deformation or rotation. | Ultrasound frames | NA | NA | NA | NA | NA | NA | 1) 2-D simulation: tissues are simulated within a region of 50 mm ×40 mm ×10 mm that lies at the depth between 50 and 100 mm, with a higher stiffness inserted at the center of the ROI; 2) 2-D phantom: tissue-mimicking gelatin phantom; a stiff cylindrical region is included and a fluid-filled thin channel is present in the phantom; 3) 3-D simulation: six scans with different tissue motions | True tissue strains are known in 2-D simulation; unknown true tissue displacements for phantom data. | NA | 1) Square root of the mean square errors (SRMSE); 2) Correlation coefficient | SRMSEs of strain estimations on 2-D simulation data: The coupled filtering method: epsilon_xx: 2.68%; epsilon_xy: 2.28%; epsilon_yx: 5.07%; epsilon_yy: 2.68%; The affine warping method: epsilon_xx: 2.82%; epsilon_xy: 2.22%; epsilon_yx: 3.80%; epsilon_yy: 2.82%.
A generalized correlation-based model for out-of-plane motion estimation in freehand ultrasound | IEEE Transactions on Medical Imaging | September 30, 2013 | Afsham et al. proposed a new out-of-plane motion estimation method that combined a closed-form derivation with a linear regression model of the ultrasound beam profile, and applies second-order statistics of the Rician-Inverse Gaussian model to improve reliability and flexibility in speckle tracking. | Ultrasound frames | NA | NA | NA | NA | NA | NA | 1) Beef tenderloin, chicken breast, turkey breast, and FDS phantom: for each tissue type, five datasets over the displacement of 4 mm (40 frames) were acquired (20 subsets in total); 2) Forearm: 75 tracked ultrasound frames from the forearm of a subject. | For data in 1): the elevation displacement was produced by means of a linear motor stage (T-LSR150B, Zaber Technologies Inc., Vancouver, BC, Canada) with equal elevation distances of 0.099 mm. The ultrasound transducer was affixed to the optical table by means of a mechanical arm. For data in 2): Optotrak Certus optical tracker (Northern Digital Inc, Waterloo, ON, Canada). | NA | 1) Accumulative estimated elevation distance versus the true value; 2) The ratio of the measured displacement error to the true elevational separation; 3) The mean target registration error (mTRE) between the true and the estimated transformation at the center of patches divided by the true elevational separation; 4) The estimated tilt angles; 5) The estimated yaw angles. | Line plots; exact numbers are not provided.
Learning for graph-based sensorless freehand 3D ultrasound | Machine Learning in Medical Imaging | October 01, 2016 | Tetrel et al. reduced error accumulation by modeling frames and motion measurements as a graph, generating random trajectories via constrained shortest paths, and refining estimates with edge weights predicted by a Gaussian process regressor. | A sequence of N frames | NA | NA | NA | NA | NA | NA | Speckle phantom: 1) Scanned elevationally over 50 mm at regular 0.05 mm intervals; 2) Nine freehand sequences; total displacements up to 30 mm. | For data 1): mounting the transducer on a robot arm following a purely translational trajectory, controlled by a step motor with a precision of 5 μm; For data 2): tracked with a Micron Tracker optical sensor. | Tested on the eight other sequences. | Mean target registration error (mTRE) of the last frame of each sequence (measured at the patch centers) with respect to the optical sensor data. | 1) The first sequence s1: 1.803; 2) s2: 0.731; 3) s3: 3.039; 4) s4: 1.995; 5) s5: 3.735; 6) s6: 9.345; 7) s7: 9.154; 8) s8: 9.999.
Deep learning for sensorless 3D freehand ultrasound imaging | Medical Image Computing and Computer-Assisted Intervention | September 04, 2017 | Among the earliest works to propose an end-to-end system for pose estimation using convolutional neural networks (Prevost et al, 2017). | Two successive images, sampled from 20 scans (phantom dataset), or from 88 scans (forearm dataset), or from 12 scans (low legs dataset), using 2-fold patient cross validation for each dataset separately. | 1) Two successive images; 2) Optical flow (two channels) | Optical flow | CNN | One transformation between two successive images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | Euclidean distance between output (6-DoF) and ground truth | No | 1) BluePhantom ultrasound biopsy phantom (20 ultrasound sweeps; 7168 frames; proximal to distal direction) 2) Forearms of 12 volunteers (88 in-vivo ultrasound sweeps; 41,869 frames; proximal to distal direction) 3) Lower legs on a subset of the volunteers (12 in-vivo ultrasound sweeps; 6647 frames; proximal to distal direction); for testing generalization capabilities. | Tracked by a surgical system (Stryker Navigation System III) | 2-fold patient cross validation for each dataset | 1) Average absolute parameter-wise error, computed per frame relative to the first frame of the sweep; 2) Final drift: the distance between the estimated and ground truth centers of the last image. | 1) Phantom dataset: avg. absolute error: t_x:1.32 mm; t_y: 2.13 mm; t_z: 7.79 mm; theta_x: 2.32 deg; theta_y: 1.21 deg; theta_z: 0.90 deg; final drift: 18.30 mm (median); 2) Forearms dataset: avg. absolute error: t_x: 3.54 mm; t_y: 3.05 mm; t_z: 4.19 mm; theta_x: 2.63 deg; theta_y: 2.52 deg; theta_z: 1.93 deg; final drift: 14.44 mm (median) ; 3) Lower legs dataset: avg. absolute error: t_x: 4.14 mm; t_y: 3.91 mm; t_z: 17.12 mm; theta_x: 1.94 deg; theta_y: 2.58 deg; theta_z: 2.15 deg; final drift: 40.56 mm (median).
A probe-camera system for 3D ultrasound image reconstruction | Imaging for Patient-Customized Simulations and Systems for Point-of-Care Ultrasound | September 08, 2017 | Ito et al. presented a probe-camera system for 3D ultrasound reconstruction, using phantom-based calibration and structure from motion for probe localisation and camera motion estimation. | Camera images | NA | NA | NA | NA | Reprojection error: Euclidean distance between a feature point and a point obtained by projecting a 3D point onto the image using a projection matrix of a camera. | No | Area around the arm and the thigh from 2 volunteers; travel distance of the US probe: 200 mm for the arm and 100 mm for the thigh. | Laser scanner (Konica Minolta, Inc., VIVID910) | NA | Root Mean Square (RMS) between the reconstructed 3D point clouds and the ground-truth mesh model | 2 mm for the probe travel distance of 200 mm
3D freehand ultrasound without external tracking using deep learning | Medical Image Analysis | August, 2018 | An extension of Prevost et al. 2017, with the following contributions: 1) integration of an IMU sensor; 2) inclusion of an additional dataset comprising 700 ultrasound sweeps with more realistic motions, including significant rotations; and 3) analysis of the impact of various acquisition parameters on reconstruction performance. | 1) Two successive images, sampled from 20 scans (phantom dataset), or from 88 scans (forearm dataset), or from 12 scans (low legs dataset), or from 600 scans (forearm dataset), or from 100 scans (carotid dataset) using 2-fold cross validation (additional pairs were generated by considering image pairs that are not consecutive); 2) IMU | 1) Two successive images; 2) Optical flow (two channels); 3) IMU (concatenate the three Euler rotation angles from the IMU with the 512-dimensional vector from the penultimate layer of the network) | Optical flow | CNN | One transformation between two successive images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | L2 loss, the squared difference between the network output and the ground truth. | No | 1) BluePhantom ultrasound biopsy phantom (20 ultrasound sweeps; from proximal to distal direction; 7168 frames; avg. length of 131 mm) 2) Forearms of 12 volunteers (88 in-vivo ultrasound sweeps, from proximal to distal direction; 41,869 frames; avg. length of 190 mm) 3) Lower legs on a subset of the 12 volunteers (12 in-vivo ultrasound sweeps, from proximal to distal direction; 6647 frames; avg. length of 175 mm); for testing generalization capabilities. 4) Forearms on another set of 15 volunteers (600 in vivo ultrasound sweeps; 307,200 frames; avg. length of 202 mm; with IMU data) 5) Carotids from 10 volunteers (100 ultrasound sweeps; 21,945 frames; avg. length of 75 mm; with IMU data); for testing generalization capabilities. | Stryker NAV3™ Camera (Stryker Co., Kalamazoo, MI, USA) optical tracking system | 2-fold cross validation | 1) Average absolute parameter-wise error: mean absolute difference between estimated and ground truth transformation parameters across all frames in each sweep; 2) Final drift of a sweep: the distance between the last frame center in the estimated trajectory and the corresponding center in the ground truth trajectory; 3) Maximum center error: the largest distance within a sweep between any estimated frame center and its corresponding ground truth center; 4) Length error: the difference in distance between the first and last frames of the estimated trajectory and that of the ground truth trajectory. | 1) Dataset # 1 (phantom): avg. absolute error: t_x: 1.32 mm; t_y: 2.13 mm; t_z: 7.79 mm; theta_x: 2.32 deg; theta_y: 1.21 deg; theta_z: 0.90 deg; final drift: 18.30 mm (median); 2) Dataset # 2 (forearms without IMU): avg. absolute error: t_x: 3.54 mm; t_y: 3.05 mm; t_z: 4.19 mm; theta_x: 2.63 deg; theta_y: 2.52 deg; theta_z: 1.93 deg; final drift: 14.44 mm (median); 3) Dataset # 3 (calves): avg. absolute error: t_x: 3.11 mm; t_y: 5.86 mm; t_z: 5.63 mm; theta_x: 2.75 deg; theta_y: 3.17 deg; theta_z: 5.24 deg; final drift: 19.69 mm (median); 4) Dataset # 4: avg. absolute error: t_x: 2.75 mm; t_y: 2.41 mm; t_z: 4.36 mm; theta_x: 0.19 deg; theta_y: 0.21 deg; theta_z: 0.13 deg; final drift: 10.42 mm (median);
Computing CNN loss and gradients for pose estimation with Riemannian Geometry | Medical Image Computing and Computer Assisted Intervention | September 26, 2018 | Hou et al. presented a Riemannian formulation for pose estimation, training CNNs on the SE(3) manifold with a left-invariant Riemannian metric and using geodesic distance as the loss to couple translation and rotation. | One image | One image | No | Any CNN architecture, e.g., GoogLeNet | The Riemannian Logarithm of canonical left-invariant metric on SE(3) | Geodesic distance between ground truth and prediction | No | 1) PoseNet dataset 2) Dataset of C-Arm X-Ray to Computed Tomography (CT) alignment problem 3) Fetal Magnetic Resonance Imaging | Not clear | Yes | For dataset 1) and 2): average errors of each parameter (Euler angles in degrees and translation in meters/millimetres) and geodesic distance; For dataset 3): Cross Correlation, Mean Squared Error, Peak Signal-to-Noise Ratio, Structural SIMilarity, and geodesic distance. | 1) Dataset 1: theta_x: 1.870 deg; theta_y: 3.143 deg; theta_z: 3.662 deg; t_x: 1.759 m; t_y: 1.240 m; t_z: 0.156 m; geodesic distance: 16.370; 2) Dataset 2: theta_x: 10.511 deg; theta_y: 6.789 deg; theta_z: 11.913 deg; t_x: 62.588 mm; t_y: 68.747 mm; t_z: 54.110 mm; geodesic distance: 19624.246; 3) Dataset 3: CC: 0.8732; MSE: 724.9713; PSNR: 20.7484; SSIM: 0.6470; geodesic distance: 10.0836
Novel similarity metric for image-based out-of-plane motion estimation in 3D ultrasound | Annual International Conference of the IEEE Engineering in Medicine and Biology Society (EMBC) | October 07, 2019 | Balakrishnan et al. proposed a similarity metric that correlated the parametric representations of image texture between consecutive ultrasound images, modeling texture dynamics with a parametrical auto-regressive (AR) model, and estimated out-of-plane motion using a trained fine Gaussian SVM regression model. | Two consecutive ultrasound images | NA | NA | NA | NA | NA | NA | Thirteen forearm ultrasound scans from three human subjects; 7503 ultrasound frames. | ElectroMagnetic Tracking System (Ascension trakSTAR) | Yes. 8 scans from subject one were used for training the SVM model and the rest were used for prediction. | 1) Parameter-wise average absolute error: the mean absolute difference between predicted and ground truth parameters across all frames in a scan; 2) Final drift: the distance between the last frame center in the prediction and in the ground truth. | 1) Parameter-wise average absolute error: t_z: 9.11 mm; theta_x: 1.95 deg; theta_y: 1.66 deg; 2) Final drift: 6.59 mm (median)
Sensorless freehand 3D ultrasound reconstruction via deep contextual learning | Medical Image Computing and Computer Assisted Intervention | September 29, 2020 | Guo et al. proposed a deep contextual network that applied 3D convolutions to ultrasound video segments with a self-attention module to emphasize speckle-rich regions and a case-wise correlation loss to stabilise training. | Sequences of 5 frames, sampled from 500 ultrasound scans for training and 70 scans for validation. | A sequence of 5 frames | No | 3D ResNext model | One transformation (mean motion vector), consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Mean squared error (MSE) loss (ground truth labels are defined as the mean parameters of the relative transform parameters between the neighboring frames); 2) Case-wise correlation loss: the Pearson correlation coefficient between the estimated motion and the ground truth mean, for every degree-of-freedom. | No | Prostate ultrasound: 640 TRUS videos (from 640 patients); varied lengths. | EM-tracking device | Yes. training: 500 cases; validation: 70 cases; testing: 70 cases. During inference, sliding along the video sequence with a window size N. The inter-frame motion of two neighboring frames is the average motion computed in all the batches. | 1) Distance error: the mean distance between corresponding corner points across all frames in the video; 2) Final drift: the distance between the center points of the transformed end frames of a video segment using the ground truth and prediction. | Distance error: 10.33 mm Final drift: 17.39 mm
Three-dimensional thyroid assessment from untracked 2D ultrasound clips | Medical Image Computing and Computer Assisted Intervention | September 29, 2020 | Wein et al. proposed a pipeline that reconstructed 3D volumes from two freehand ultrasound sweeps (transverse and sagittal) using deep‑learning trajectory estimation and image‑based 3D model optimisation. | CNN: two successive ultrasound frames, sampled from 100 sweeps during training; Registration (optimision): transversal and sagittal sweeps, reconstructed from CNN. | CNN: two successive ultrasound frames; two channels of 2D optical flow; Registration (optimision): transversal and sagittal sweeps, reconstructed from CNN. | CNN: two channels of 2D optical flow | CNN | CNN: one transformation between two successive images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) CNN: Euclidean distance between output (6-DoF) and ground truth; 2) Registration: image similarity on the B-Mode intensities (optimising a number of trajectory parameters and the rigid transformation between the sweep centers). | No | Thyroid data: 180 ultrasound sweeps from 9 volunteers; 5 transversal and 5 sagittal sweeps for each lobe of the thyroid. | Tracked using an optical tracking system; Manually segmented on 58 sweeps (50 from the 5 volunteer training subset and 8 from the 4 volunteer testing subset.) | Yes. Training: 100 sweeps from 5 volunteers; testing: 80 sweeps from 4 volunteers. | 1) Dice coefficient: the overlap between the 3D thyroid gland segmentations derived from the transversal and sagittal sweeps; 2) Relative trajectory error: the cumulative in-plane translation error at each frame divided by the ground truth sweep length; 3) Relative length error: the ratio of the difference between the estimated and ground truth sweep lengths to the ground truth length; 4) Volume and Norm. vol. | Results from 54-DOF refinement: 1) Dice: 0.66±0.12; 2) Relative trajectory error: 0.16±0.09 3) Relative length error: 0.16±0.14 4) Volume: 1.15±1.45; Norm. vol.: 0.14±0.16
Localizing 2D ultrasound probe from ultrasound image sequences using deep learning for volume reconstruction | Medical Ultrasound, and Preterm, Perinatal and Paediatric Image Analysis | October 01, 2020 | Miura et al. presented a CNN with feature extraction and motion estimation components, along with a consistency loss to enforce. | Two successive ultrasound frames, sampled from 180 scans (27,948 frames) for training and 30 scans (5,176 frames) for validation. | Two successive ultrasound frames | Extracted optical flow using FlowNetS | Localization network: ResNet34; Optical flow estimation network: FlowNetS (pre-trained by the Flying Chairs dataset) | One transformation between two successive images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Euclidean distance between ground truth and the estimated 6 parameters; 2) Reprojection error (left-right consistency loss): Euclidean distance between a point on the image and a point reprojected from the consecutive image frame. | No | 1) Forearms from 5 subjects: 190 scans (30,801 frames) 2) Breast phantom: 60 scans (8,940 frames) 3) Hypogastric phantom: 40 scans (6,242 frames) | V120:Trio (OptiTrack) | Training data: 180 scans (27,948 frames) from forearm of 2 subjects and two phantoms; validation data: 30 scans (5,176 frames) from forearm of 1 subject; test data: 80 scans (12,859 frames) from forearm of 2 subjects. | Mean absolute error (MAE) for each parameter | MAE: theta_x: 0.53 deg; theta_y: 1.21 deg; theta_z: 0.47 deg; t_x: 0.64 mm; t_y: 0.15 mm; t_z: 0.80 mm
Probe localization from ultrasound image sequences using deep learning for volume reconstruction | Proc. SPIE 11792, International Forum on Medical Imaging in Asia | April 20, 2021 | Miura et al. estimated in-plane and out-of-plane probe motion using CNN, and introduced two loss functions to guarantee the consistency of forward and backward motion of the probe. | Two successive ultrasound images, sampled from 370 ultrasound scans for training and 40 scans for validation. | 1) Two successive ultrasound images; 2) Optical flow | Optical flow | AlexNet | One transformation between two successive images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Euclidean distance between prediction and ground truth; 2) Forward consistency loss: Euclidean distance between points projected forward through successive frame-to-frame transformations and the ground truth points. 3) Backward consistency loss: Euclidean distance between the original points and those reprojected forward and then back through successive frame-to-frame transformations. | No | 1) Forearm (310 scans); 2) Breast phantom (180 scans) 3) Hypogastric phantom (40 scans) | OptiTrack V120: Trio (Acuity Inc.) | Training: 370 scans from the forearms of 3 subjects and the breast phantom; validation: 40 scans from the forearm of 1 subject; test: 120 scans from the forearms of 2 subjects and the hypogastric phantom. | 1) Mean absolute error (MAE): the average of cumulative errors from the initial frame; 2) Final drift: the error in the estimated probe position of the final frame. | 1) Forearm data: MAE: theta_x: 13.6 deg; theta_y: 18.9 deg; theta_z: 7.5 deg; t_x: 8.5 mm; t_y: 4.3 mm; t_z: 15.0 mm; final drift: 75.8 mm; 2) Hypogastric phantom: MAE: theta_x: 9.7 deg; theta_y: 16.0 deg; theta_z: 11.3 deg; t_x: 12.9 mm; t_y: 9.9 mm; t_z: 30.5 mm; final drift: 73.6 mm.
Transducer adaptive ultrasound volume reconstruction | IEEE International Symposium on Biomedical Imaging | May 25, 2021 | Guo et al. addressed transducer-related domain shifts by proposing a domain adaptation strategy that trained feature extractors to produce transducer-invariant features by minimising latent-space discrepancies between paired samples. | Sequences of 5 consecutive frames, sampled from 500 source domain scans for training, 70 source domain scans for validation, and 9 target domain scans for training. | A sequence of 5 consecutive frames | No | CNN | One transformation (mean motion vector), consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | Mean squared error (MSE) loss between network output and ground truth DOF labels. | Discrepancy loss: the L2 norm between the two feature extractors’ output feature vectors, from source domain data and target domain data. | 1) Source domain: 640 prostate transrectal ultrasound video sequences; 2) Target domain: 12 transabdominal ultrasound video sequences. | EM-tracking device | Source domain: training: 500 cases; validation: 70 cases; testing: 70 cases; Target domain: training: 9 cases; testing: 3 cases; During inference, using a sliding-window procedure to get the DOF motion vector prediction for every subsequence. | 1) Distance error; 2) Final drift. | On target domain: Distance error: 12.67 mm; Final drift: 20.34 mm.
Learning to map 2D ultrasound images into 3D space with minimal human annotation | Medical Image Analysis | May 2021 | Yeung et al. proposed a self-supervised model that learns from 2D slices of 3D fetal brain volumes by pairwise attention-based comparisons, capturing relative positional information for prediction. | Sets of an arbitrary number of 2D images, sampled from 3D ultrasound fetal brain volumes. | A set of an arbitrary number of 2D images | No | CNN + attention mechanism | The set of corresponding predicted locations (3 Cartesian coordinates (i.e. x,y,z) of the centre, the bottom right and left corners of a plane) | Mean squared error loss between the ground truth and predicted locations | No | 3D ultrasound fetal brain volumes, obtained as part of the INTERGROWTH-21st study; 18 to 22 gestational weeks; fifty and fifteen 3D volumes acquired at 21 gestational weeks were selected for training and validation; 150,000 and 45,000 images for each training and validation epoch. | Artificially sample 2D slices from aligned 3D ultrasound volumes of the fetal brain | 1) Trained on images at 21 gestational weeks, testing on 18 to 22 gestational weeks; 2) Training and testing on 18 to 22 gestational weeks; 3) Testing on real 2D images taken at the standard TT plane; 4) Testing on 15 videos of 2D brain scans acquired from 4 subjects with gestational age between 19 to 21 weeks. | 1) Euclidean distance (ED) between all the coordinates of the predicted and ground truth planes; 2) Plane angle (PA) between the predicted and ground truth planes; 3) Normalised cross-correlation (NCC) between the input image and image sampled from the predicted plane. | Testing on real 2D images taken at the standard TT plane: 1) Comparison with the first manual annotation: ED: 11.36±3.26 voxels; PA: 0.179±0.095 rad; NCC: 0.841±0.096; 2) Comparison with the second manual annotation: ED: 11.44±5.02 voxels; PA: 0.180±0.120 rad; NCC: 0.837±0.080.
Self context and shape prior for sensorless freehand 3D ultrasound reconstruction | Medical Image Computing and Computer Assisted Intervention | September 21, 2021 | Luo et al. proposed a recurrent convolutional LSTM based online learning framework via a differentiable reconstruction algorithm, a self‑supervised learning method that exploited contextual cues, and adversarial training for anatomical shape prior learning, improving robustness to complex scan sequences. | N-length scans, sampled from 8500 sequences generated from 85 3D DDH scans for training, or 6500 sequences generated from 65 3D fetus ultrasound scans for training. | N-length scan | No | Convolutional LSTM (ConvLSTM) | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error (MAE) loss between prediction and ground truth; 2) Case-wise correlation loss; 3) A cross-entropy loss that trains the classifier to distinguish reconstructed volumes from real volumes. Test phase: Self-supervised loss between the original frames and the corresponding slices extracted from the reconstructed volume at the estimated positions. | Test phase: 1) Adversarial loss; 2) Quadratic potential divergence. | 1) 3D developmental dysplasia of the hip (DDH) ultrasound dataset from 14 volunteers (101 US volumes) 2) 3D fetus ultrasound dataset from 78 volunteers (78 fetal ultrasound volumes); gestational age ranges from 10 to 14 weeks. | Complex ultrasound scan were simulated from real 3D volumes by combining loop, fast-and-slow scan, and sector scans. The number of generated sequences is 100 and 10 for training and testing, respectively. | Yes. 1) 3D DDH dataset: training: 85 volumes; testing: 15 volumes; 2) Fetal US dataset: training: 65 volumes; testing: 13 volumes. | 1) Final drift rate (FDR): the final drift divided by the scan length; 2) Average drift rate (ADR): the mean value of the cumulative drift of all frames divided by the length from the frame to the starting point of the sequence; 3) Maximum drift (MD): the maximum accumulated drift of all frames; 4) Sum of drift (SD): the sum of accumulated drift of all frames; 5) Bidirectional Hausdorff distance (HD): the worst distances between the predicted positions and the ground truth of all the frames in the sequence. | DDH dataset: 1) FDR: 5.44±3.03 %; 2) ADR: 13.47±6.47 %; 3) MD: 4.44±2.29 mm; 4) SD: 274.14±136.89 mm; 5) HD: 3.91±1.89 mm; Fetus dataset: 1) FDR: 9.94±4.41 %; 2) ADR: 27.08±9.31 %; 3) MD: 16.84±7.56 mm; 4) SD: 730.11±303.20 mm; 5) HD: 14.12±5.89 mm.
Pose estimation of 2D ultrasound probe from ultrasound image sequences using CNN and RNN | Simplifying Medical Ultrasound | September 21, 2021 | Miura et al. integrated CNN and RNN to exploit long‑term temporal context, predicting both relative and absolute probe poses directly from image sequences. | Sequences of 11 frames, sampled from 300 scans for training and 30 scans for validation. | 1) CNN (feature extraction): two consecutive frames and two channels of optical flow; 2) ConvLSTM: one feature map from CNN. | Two channels of optical flow (estimated by FlowNetS, trained on Flying Chairs dataset) | CNN (AlexNet; trained on ImageNet) + ConvLSTM + spatial-temporal attention module | 10 transformations between adjacent frames and 11 transformations (absolute pose); each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | For relative pose estimation: 1) Euclidean distance between prediction and ground truth; 2) Forward consistency loss: Euclidean distance between points projected forward through successive frame-to-frame transformations and the ground truth points; 3) Backward consistency loss: Euclidean distance between the original points and those reprojected forward and then back through successive frame-to-frame transformations. For absolute pose estimation: MSE between ground truth pose and estimated absolute pose | No | 1) Forearms from 6 subjects (230 scans) 2) Breast phantom (140 scans) 3) Hypogastric phantom (40 scans) | OptiTrack V120: Trio (Acuity, Inc.) | Yes. Experiment 1): Only use breast phantom data with straight line scans: training: 40 scans; validation: 5 scans; testing: 5 scans; Experiment 2): All three datasets: training: 300 scans; validation: 30 scans; testing: 80 scans. | 1) Sum of absolute errors (AE) of the relative pose parameters per frame, computed separately for each parameter; 2) Final drift: the distance between the estimated and ground truth positions of the final frame. | 1) Breast phantom: AE: theta_x: 19.52 deg; theta_y: 26.38 deg; theta_z: 22.04 deg; t_x: 23.97 mm; t_y: 14.02 mm; t_z: 45.87 mm; final drift: 27.71 mm; 2) On three datasets: AE: theta_x: 125.37 deg; theta_y: 252.70 deg; theta_z: 120.18 deg; t_x: 133.24 mm; t_y: 40.63 mm; t_z: 184.08 mm; final drift: 96.46 mm.
ImplicitVol: Sensorless 3D ultrasound reconstruction with deep implicit representation | arXiv | September 24, 2021 | Yeung et al. introduced a model for sensorless 3D reconstruction from 2D freehand ultrasound scans. Instead of voxel grids, it represented the 3D volume implicitly as a continuous function, while jointly refining scan locations. | 3D locations of N images, computed from estimated transformations (3 for Euler angle, 3 for translation); 2D images are sampled from fifteen 3D volumes around the central axis of the brain non-uniformly, or sampled from two videos of real 2D freehand fetal brain ultrasound; N = 128 or 256 | 3D locations of N images, computed from estimated transformations (3 for Euler angle, 3 for translation); N = 128 or 256 | Positional encoding | MLP | Intensities of pixels of N input frames | Structural similarity (SSIM) loss between the actual and network-predicted intensities | No | 1) Volume-sampled 2D images: sampling planes from fifteen 3D ultrasound fetal brain volumes; 20 weeks gestational age; obtained as part of the INTERGROWTH-21st study; 2) Two videos of native freehand 2D brain scans: 250 frames each; 20 weeks gestational age. | 1) For volume-sampled 2D images: ground truth is from the original volumes; 2) For freehand 2D brain scans: lack of ground truth 3D location information | Yes | 1) Normalised cross-correlation (NCC); 2) Structural similarity index measure (SSIM); 3) Absolute difference between rotation angles; 4) Absolute distance between translations. | On volume-sampled 2D images (N=256): 1) Slices sampled at axial directions: NCC: 0.747±0.096; SSIM: 0.569±0.085; 2) Slices sampled at coronal directions: NCC: 0.751±0.098; SSIM: 0.570±0.080; 3) Slices sampled at sagittal directions: NCC: 0.701±0.116; SSIM: 0.606±0.072; 4) Location: Angle: 0.137±0.097 rad; Distance: 6.32±4.61 pixel
Image-based 3D ultrasound reconstruction with optical flow via pyramid warping network | IEEE Engineering in Medicine & Biology Society | December 09, 2021 | Xie et al. fused image and optical flow features via a pyramid warping layer with channel and spatial attention. | Two adjacent ultrasound images, sampled from 100 ultrasound sweeps using 10-fold cross validation. | 1) Two adjacent ultrasound images; 2) Optical flow. | Optical flow with 4 different scales, predicted from FlowNetCSS module | ResNet + channel attention + FlowNetCSS + spatial attention | One transformation between two adjacent images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Mean absolute error (MAE) loss: distance between predicted 6-DoF and ground truth; 2) Transformation matrix loss: the Frobenius norm between 4×4 transformation matrixes converted from 6-DoF of output and ground truth. | No | Forearms of 100 ultrasound sweeps (19502 frames) | NDI Polaris Vicra | 10-fold cross validation | 1) Mean distance error: average distance between the predicted and ground truth positions of all frames within a sweep; 2) Final drift: the distance of the final frame. | 1) Distance error: 4.73 mm; 2) Final drift: 8.55 mm.
Spatial position estimation method for 3D ultrasound reconstruction based on hybrid transfomers | IEEE International Symposium on Biomedical Imaging | April 26, 2022 | Ning et al. leveraged transformer for performing regression tasks on the sequence, within a joint local and global information encoding approach. | 1) Sequences of N frames, sampled from 60 cases for training and 20 cases for validation (forearm dataset), or from 96 cases for training and 32 cases for validation (clinical dataset); 2) Quaternions of N frames from IMU data. | 1) A sequence of N frames; 2) Quaternions of N frames from IMU data. | 1) Position embedding; 2) Difference map. | CNN + Transformer | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | MSE between the ground truth of relative poses and predictions | No | 1) Forearm dataset: 100 cases (200 frames each); absolute range of probe movement: 10-20 cm; 2) Clinical dataset: 160 cases (200 frames each); absolute range of probe movement: 15-25 cm. | NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. The dataset was split into training, validation, and testing sets in a 3:1:1 ratio. | 1) Final drift: the distance between the center point of the first and last frame in the label and the distance being predicted; 2) Average distance error: the mean distance between corresponding corner points of each frame pair across the entire sequence. | 1) Forearm dataset: final drift: 15.84 mm; average distance error: 8.99 mm; 2) Clinical dataset: final drift: 35.17 mm; average distance error: 27.91 mm
Stretched reconstruction based on 2D freehand ultrasound for peripheral artery imaging | International Journal of Computer Assisted Radiology and Surgery | April 29, 2022 | Leblanc et al. used Mask-RCNN segmentation, in-plane registration, and CNN-predicted out-of-plane translation to generate a 3D stretched reconstruction of the femoral artery. | Two consecutives frames, sampled from 13,087~15,687 consecutives pairs of original frames for training and 4323~4916 consecutives pairs of original frames for validation. | Two consecutives frames | No | Mask R-CNN with inception backbone (for segmenting the femoral artery) + CNN | Relative longitudinal motion (in mm) between the input registered pair of frames | L1 loss between ground truth and prediction | No | Left or right femoral artery from 18 healthy volunteers: 111 ultrasound sweeps; 40,788 2D frames; proximal to distal. | 1) For pose: optical tracking device (NDI Polaris Spectra); 2) For segmentation: manually annotated using LabelMe; manually segmented (12 out of the 111 sequences from 6 volunteers) | For pose: 5-fold cross-validation. Training: 10 volunteers (13,087~15,687 consecutives pairs of original frames); validation: 4 volunteers (4323~4916 pairs); testing: 4 volunteers (4058~5789 pairs). For segmentation: training: 6 sequences (4 volunteers, 2073 frames); validation: 3 sequences (1 volunteer, 702 frames); testing: 3 sequences (1 volunteer, 1004 frames); | 1) Dice score between the manual segmentation and the output masks; 2) Relative frame-to-frame longitudinal error: the absolute difference between the prediction and ground truth; 3) Final drift: the absolute difference between the position of the last frame reconstructed using tracker-provided ground truth and the position estimated from the network outputs. | 1) Dice score: a false negative rate of 7% and a mean dice score of 0.90 computed on the true positive cases; 2) Relative frame-to-frame longitudinal error: 0.28 mm; 3) Final drift: 17.22 mm
Deep learning-based plane pose regression in obstetric ultrasound | International Journal of Computer Assisted Radiology and Surgery | April 30, 2022 | Di Vece et al. used a regression CNN to estimate the ultrasound plane poses in obstetric imaging. | Ultrasound images, sampled from 75088 images for training (phantom data) or 22029 images for training (real fetal data). | An ultrasound image | No | ResNet-18 | A vector of nine parameters relative to the centre of the fetal brain ultrasound volume (3 for translation and 6 for rotation); the last six parameters are used internally by the CNN to reconstruct the rotation matrix using Gram-Schmidt process. | Mean squared error (MSE) between prediction and ground truth, on translation and rotation matrices separately. | No | 1) Phantom fetal data: six brain volumes; 23-week GA fetus US examination phantom; 18047 planes with random orientation per volume and 725 around the TV SP; 2) Real fetal data: seven brain volumes; GA ranging from 21 to 25 and 39 weeks; 20699 images at random coordinates and 1330 around the TV SP for each volume. | The slices were generated by applying rotation and translation to a plane starting at the center of the volume, with parameters drawn from a uniform random distribution within a fixed range. The position of the TV SP were annotated by a sonographer by directly manipulating a slicing plane within Unity and chose the translation and angle sampling intervals. | Experiment 1 (only use phantom data): Training: 75088 images from 4 volumes; testing: 37544 images from 2 volumes; Experiment 2 (initialisation with weights from the phantom; training and testing on real data): training: one fetus, 22029 images, GA of 23 weeks; testing: six fetus ranging from a GA of 21 to 39 weeks. | 1) Euclidean distance between the two planes for calculating translation errors (in mm); 2) Geodesic distance for calculating rotation error (in degree) | 1) Experiment 1: Test planes acquired at random coordinates (Test RP): translation norm: 0.90 mm (median); rotation geodesic distance: 1.17 deg (median). Test planes acquired around TV SP (Test SP): translation norm: 0.44 mm (median); rotation geodesic distance: 1.21 deg (median). 2) Experiment 2 (Test RP): GA 21: translation norm: 9.94 mm (median); rotation geodesic distance: 30.58 deg (median). GA 22: translation norm: 10.74 mm (median); rotation geodesic distance: 30.81 deg (median). GA 23: translation norm: 10.39 mm (median); rotation geodesic distance: 21.94 deg (median). GA 24: translation norm: 17.76 mm (median); rotation geodesic distance: 37.93 deg (median). GA 25: translation norm: 18.80 mm (median); rotation geodesic distance: 42.23 deg (median). GA 39: translation norm: 17.15 mm (median); rotation geodesic distance: 34.43 deg (median).
Ultrasound volume reconstruction from freehand scans without tracking | IEEE Transactions on Biomedical Engineering | September 14, 2022 | Guo et al. expanded their work (Guo et al. 2020) by introducing the contrastive learning strategy and investigating the impact of experimental settings on reconstruction performance. | Sequences of 5 frames, sampled from 488 scans for training and 66 scans for validation, or 80 scans for training and 10 scans for validation using 10-fold cross-validation. | A sequence of 5 frames | No | 3D ResNeXt + self-attention block | One transformation (mean motion vector), consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Mean squared error (MSE) loss between prediction and ground truth (Ground-truth labels are defined as the mean parameters of the relative transform parameters between the neighboring frames); 2) Case-wise correlation loss: the Pearson correlation coefficient between the estimated motion and the ground truth mean, for every degrees-of-freedom. | Margin ranking loss: correlates two random samples' feature vectors distance with their transformation parameters discrepancy. | 1) Dataset A (Prostate): 618 transrectal ultrasound video sequences; 618 patients; 2) Dataset B (Prostate): 100 transabdominal/transperineal ultrasound video sequences; <100 patients; varying lengths and resolutions. | EM tracking device | Yes Dataset A: training: 488 cases; validation: 66 cases; testing: 64 cases; Dataset B: training: 80 cases; validation: 10 cases; testing: 10 cases; 10-fold cross-validation. | 1) Distance Error: the average distance between corresponding corner points of the input patches across the sweep; 2) Frame Error: the difference between the ground truth and predicted frame locations of neighboring frames; 3) Final Drift: the Euclidean distance between the last frame of the ultrasound video; 4) Drift Rate: the ratio between the final drift and the ground truth sequence length; 5) Prostate Dice: the Dice coefficient between the prostate’s ground truth segmentation and the predicted segmentation; 6) Prostate Error: the absolute volume difference between the ground truth segmentation and the predicted segmentation in cubic centimeters. | Dataset A: Distance Error: 5.52±2.86 mm; Frame Error: 0.90±0.26 mm; Final Drift: 10.20±8.47 mm; Drift Rate: 9.64±8.14 %; Prostate Dice: 0.89±0.06; Prostate Error: 3.21±1.93 cc; Dataset B: Distance Error: 5.87±2.68 mm; Frame Error: 1.12±0.26 mm; Final Drift: 9.85±5.74 mm; Drift Rate: 14.58±12.76 %;
Deep motion network for freehand 3D ultrasound reconstruction | Medical Image Computing and Computer Assisted Intervention | September 16, 2022 | Luo et al. introduced IMU for estimating elevational displacements outside the plane with an online self-supervised strategy. | 1) N-length scans, sampled from 196 scans for training (arm dataset) or from 136 scans for training (carotid dataset) (each scan was randomly augmented to 40 sequences at each epoch); 2) Acceleration from IMU; 3) Euler angle calculated using orientation data from IMU. | 1) A N-length scan; 2) Acceleration from IMU; 3) Euler angle calculated using orientation data from IMU. | No | ResNet + LSTM | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error (MAE) between the estimated transformation parameters and ground truth; 2) Pearson correlation loss; Test phase: 1) Pearson correlation loss between the estimated acceleration and acceleration from IMU; 2) MAE loss between the estimated Euler angle and IMU Euler angle. | No | 1) Arm dataset: from 41 volunteers (250 scans); average length of 94.83 mm; linear, curved, fast-and-slow and loop scans; 2) Carotid dataset: from 40 volunteers (160 scans); average length of 53.71 mm; linear scans. | An electromagnetic (EM) positioning transmitter/receiver | Yes. Arm data: taining: 196 scans; test: 54 scans; Carotid data: training: 136 scans; test: 24 cans. Each scan in training set was randomly augmented (subsequence intercepting, interval sampling, and sequence inversion) to 40 sequences and regenerated at each epoch; Each scan in test set was randomly augmented to 10 fixed sequences. | 1) Final drift rate (FDR); 2) Average drift rate (ADR); 3) Maximum drift (MD); 4) Sum of drift (SD); 5) Symmetric Hausdorff distance (HD); 6) Mean error of angle (EA) | 1) Arm dataset: FDR = 12.75±9.05 %; ADR = 19.05±11.46 %; MD = 10.24±7.36 mm; SD = 332.29±316.36 mm; HD = 9.40±7.13 mm; EA = 1.55±1.46 deg; 2) Carotid dataset: FDR = 15.67±8.37 %; ADR = 25.08±9.34 %; MD = 8.89±4.31 mm; SD = 258.83±208.12 mm; HD = 8.28±4.29 mm; EA = 1.50±0.98 deg;
Trajectory estimation of ultrasound images based on convolutional neural network | Biomedical Signal Processing and Control | September 2022 | Mikaeili et al. leveraged both a densely connected network and FlowNet, to estimate the position of ultrasound images. | Sequences of 3 or 5 frames, sampled from 2220 frames during training. | A sequence of 3 or 5 frames | No | DenseNet + FlowNet | Transformation between two adjacent images, consisting of 6 degree of freedoms (3 for translation, 3 for rotation) | Not clear | L2 regularization (the squared magnitude of all weights in the network) | Model 539 Multi-purpose phantom; three sweeps (2520 frames in total) | DFRobot manufacturer’s inertial measurement unit (IMU) - Bosch BNO055 IMU | Yes. The last 100 frames from each of the three sets were held out for testing. | Average absolute error of 6 DoF parameters, between the prediction and the ground truth. | Results of input of 3 images: Data set 1: theta_x: 0.6067 deg; theta_y: 0.8364 deg; theta_z: 0.7604 deg; t_x: 2.2955 mm; t_y: 1.4807 mm; t_z: 2.3093 mm; Data set 2: theta_x: 0.4309 deg; theta_y: 0.7950 deg; theta_z: 0.7449 deg; t_x: 2.3119 mm; t_y: 1.5365 mm; t_z: 2.3676 mm; Data set 3: theta_x: 0.4193 deg; theta_y: 0.8120 deg; theta_z: 0.7434 deg; t_x: 2.3166 mm; t_y: 1.4754 mm; t_z: 2.2971 mm;
Adaptive 3D localization of 2D freehand ultrasound brain images | Medical Image Computing and Computer Assisted Intervention | September 16, 2022 | Yeung et al. proposed a sensorless framework for localising 2D ultrasound images in a 3D anatomical atlas, trained on co-aligned 3D volumes and fine-tuned on freehand scans using an unsupervised cycle-consistency constraint. | A set of n 2D ultrasound slices, randomly sampled from 50 3D ultrasound volumes which are co-aligned with the predefined 3D anatomical atlas. Fine-tuned on m volume-sampled 2D images or native 2D freehand images. | A set of n sampled 2D slices from 3D volumes for training; a new set of m 2D ultrasound images for fine-tuning. | No | CNN | 1) Locations (x, y and z coordinates of the top right, top left and bottom right corners) of input slices, in a predefined 3D anatomical atlas; 2) Displacement between each pair of slices. | 1) Mean least-squared error (MSE) between the predicted and ground truth locations; 2) MSE between the predicted and ground truth displacements; 3) Cycle consistency loss: MSE between the overall displacement of a sequence of images and the displacement from the first image to the last of that sequence. | No | 1) Volume-sampled images: 50 training and 17 testing 3D ultrasound volumes co-aligned with the predefined 3D anatomical atlas; 3000 slices for each testing volume; 2) Native 2D freehand images: video sequences of 2D freehand ultrasound brain scans. | 1) Volume-sampled images: sampled from aligned 3D ultrasound volumes; 2) Native 2D freehand images: the ground-truth locations were not available. | No. Consider each ultrasound acquisition from different machines as a different domain. Training data and test data are from different domains. Training: volume-sampled 2D images, from 50 3D volumes; fine-tuning and inference: volume-sampled 2D images from 17 aligned 3D volumes and native 2D freehand images. | 1) Euclidean distance (ED) between the coordinates of the predicted and ground truth planes in the predefined 3D anatomical atlas; 2) Dihedral angle (DA) between the predicted and ground truth planes; 3) Normalised standard deviation of change rate: quantifies the smoothness of consecutive ultrasound image transitions, with lower values indicating greater consistency. | For volume-sampled images: 1) Fine-tune and test on the same set of images: ED: 23.7±9.01 voxel; DA: 0.198±0.092 rad; 2) Fine-tune and test on different set of images: ED: 33.0±15.1 voxel; DA: 0.211±0.097 rad. For native freehand images: Normalised standard deviation of change rate: 0.553
A freehand 3D ultrasound reconstruction method based on deep learning | Electronics | March 23, 2023 | Chen et al. employed a 3D CNN–LSTM to estimate pose, using as input both original ultrasound frames and frames generated through Bezier interpolation and speckle decorrelation. | Three original images sampled from one scan, and eight interpolated images. | Three original images and eight interpolated images | Bezier interpolation among input ultrasound frames, to make the FPS is consistent with the tracking device. | 3DCNN + LSTM | 1) Position of ROI middle pixel; 2) Six offset parameters. | Mean square error loss | No | 1) 1 mm metal sheet: 1000 frames; average distance of 15 mm; 2) 2 mm metal sheet: 1000 frames; average distance of 16 mm; 3) Symmetrical metal block: 2000 frames; average distance of 26 mm; 4) Asymmetric metal block: 2000 frames; average distance of 21 mm; 5) Right of human arm: 1500 frames; average distance of 100 mm; 6) Left of human arm: 1500 frames; average distance of 100 mm. | Acousto-optic positioning equipment | Yes. Each of the six datasets was split into a training set and a test set at a ratio of 7:3. | 1) Theta X; 2) Theta Y; 3) Theta Z; 4) Lateral X; 5) Lateral Y; 6) Final drift | 1) Dataset 1): X: 1.55 mm; Y: 1.55 mm; Theta X: 0.61 deg; Theta Y: 0.67 deg; Theta Z: 0.83 deg; Final drift: 1.25 mm (median); 2) Dataset 3): X: 1.53 mm; Y: 1.55 mm; Theta X: 0.75 deg; Theta Y: 0.75 deg; Theta Z: 1.03 deg; Final drift: 1.33 mm (median); 3) Dataset 5): X: 1.55 mm; Y: 1.55 mm; Theta X: 1.16 deg; Theta Y: 1.05 deg; Theta Z: 1.44 deg; Final drift: 1.3 mm (median); 4) Dataset 6): X: 1.55 mm; Y: 1.53 mm; Theta X: 1.1 deg; Theta Y: 1.07 deg; Theta Z: 1.49 deg; Final drift: 1.31 mm (median).
Ultra-NeRF: Neural radiance fields for ultrasound imaging | Medical Imaging with Deep Learning | April, 04 2023 | Wysocki et al. proposed a physics-enhanced implicit neural representation for ultrasound imaging that leveraged a ray-tracing-based neural rendering, enabling more faithful synthesis of geometrically accurate B-mode images from overlapping sweeps. | Positions of points in a Cartesian coordinate space, sampled along a ray from four tilted scans for training (synthetic data), or four tilted scan pairs for training (phantom data). | Positions of points in a Cartesian coordinate space, sampled along a ray | No | MLP + ultrasound volume rendering model | Parameter vector (attenuation, reflectance, border probability, scattering density, and scattering intensity) from MLP; pixel intensity from rendering. | 1) Structural Similarity Index Measure (SSIM) between the generated B-mode images with their corresponding target images; 2) Squared Error Loss (L2) | No | 1) Synthetic data: simulated B-mode images of a liver from CT images using ImFusion; seven scans; 200 2D ultrasound images for each scan; 2) Phantom data: lumbar spine, gelatine-based phantom; 13 scans; 150 frames for each scan. | 1) Not clear for synthetic data; 2) Robotic tracking for phantom data. | 1) For synthetic data: four tilted scans for training, three scans of one perpendicular and two tilted for testing. 2) For phantom data: four tilted scan pairs for training, two tilted scan pairs and one perpendicular scan for testing. The visible tissue are covered both in training and test sets. | SSIM between synthetic and reference testing data | SSIM between synthetic and reference B-mode images: 1) Liver synthetic data: tilted scan: 0.45; perpendicular scan: 0.49; 2) Spine phantom data: tilted scan: 0.51; perpendicular scan: 0.54.
RecON: Online learning for sensorless freehand 3D ultrasound reconstruction | Medical Image Analysis | July 2023 | An extension of Luo et al., 2021, by introducing path-level supervision and a motion‑weighted training loss. | N-length 2D ultrasound scans, sampled from 13500 generated scans from 135 3D DDH volumes for training, 16 volumes for validation, or 9800 generated scans from 98 3D fetus ultrasound volumes, 15 volumes for validation, or 45 2D spine ultrasound scans for training, 12 scans for validation. | N-length 2D ultrasound scan (features for each pair of adjacent frames from ResNet are fed into LSTM) | 1) Canny edge map for each frame; 2) Optical flow from the adjacent frames. | ResNet + LSTM | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error (MAE) between predicted 6-DoF and ground truth of adjacent frames; 2) Pearson correlation; 3) Cross-entropy loss that trains the classifier to distinguish reconstructed volumes from real volumes. Test phase: 1) Frame-level contextual consistency: MAE between the original frames and the corresponding slices extracted from the reconstructed volume at the estimated positions; 2) Path-level similarity constraint that selects the prior training scan whose ground truth transformations are most similar to those of the current test scan and optimises the network by minimising the MAE and Pearson correlation loss between the estimated transformations of the selected training scan and its ground truth. | Training phase: Motion-weighted regularization that builds a positive correlation between motion and appearance variance. Test phase: Global adversarial shape prior identifies the reconstructed volume from numerous real volumes in the training dataset. | 1) 3D developmental dysplasia of the hip (DDH) ultrasound dataset from 28 volunteers: 169 ultrasound volumes; 2) 3D fetus ultrasound dataset from 128 pregnant volunteers: 128 ultrasound volumes; gestational age ranges from 10 to 14 weeks; 3) 2D spine ultrasound dataset from 23 volunteers: 68 scans; mean length of 186 mm. | For datasets 1) and 2): a motion-based simulation method for generating 2D scans with complex scanning strategies from volumetric regions of interest; For dataset 3): a robotic arm with EM positioning | Yes. 1) For 3D DDH dataset: training: 135 volumes; validation: 16 volumes; test: 18 volumes; 2) For fetus US dataset: training: 98 volumes; validation: 15 volumes; test: 15 volumes; Generating 100 different scans for each volume in each training epoch. In the testing phase, generating 20 fixed scans for each volume. 3) For 2D spine US dataset: training: 45 scans; validation: 12 scans; test: 11 scan; | 1) Final Drift Rate (FDR): the final drift divided by the scan length; 2) Average Drift Rate (ADR): in which drift rate refers to the cumulative drift of a frame divided by the length from this frame to the starting frame of the scan; 3) Maximum Drift (MD): the maximum drift of all frames; 4) Sum of Drift (SD): the sum of all drift; 5) Symmetric Hausdorff Distance (HD): the maximum distance between the estimated and ground truth positions across all frames; 6) Mean Error of Angle (MEA): MAE between estimated and real Euler angles. | 1) DDH: FDR: 10.68±11.21 %; ADR: 10.39±6.71 %; MD: 5.25±5.17 mm; SD: 71.95±75.33 mm; HD: 3.09±1.87 mm; MEA: 1.33±2.02 deg; 2) Fetus: FDR: 13.31±10.94 %; ADR: 18.42±11.78 %; MD: 18.69±12.78 mm; SD: 236.37±256.35 mm; HD: 9.60±6.45 mm; MEA: 2.26±2.66 deg; 3) Spine: FDR: 10.82±5.36 %; ADR: 19.72±12.83 %; MD: 21.05±10.82 mm; SD: 660.85±909.42 mm; HD: 20.38±11.07 mm; MEA: 1.25±0.46 deg.
Trackerless freehand ultrasound with sequence modelling and auxiliary transformation over past and future frames | IEEE International Symposium on Biomedical Imaging | September 01, 2023 | To investigate the influence of past and future frames in freehand ultrasound reconstruction, Li et al. proposed a multi-task learning algorithm which utilised a large number of auxiliary transformation-predicting tasks. | Sequences of 20 frames, sampled from 137 scans for training and 46 scans for validation. | A sequence of 20 frames for EfficientNet (b1) or one frame for LSTM | No | EfficientNet (b1) or EfficientNet + LSTM | 79 selected transformations from C_20^2 transformations; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Mean-square-errors (MSEs) between four corner points, transformed from predictions and ground truth, average over all predictions. 2) Accumulated loss: MSE between four corner points obtained from the ground truth transformation and from an indirect prediction composed of two direct predictions. | Consistency loss: MSE between four corner points obtained from a direct prediction (e.g., transformation from frame 1 to frame 10) and from an indirect prediction composed of two direct predictions (e.g., frame 1 to frame 5, then frame 5 to frame 10) | Forearms from 19 volunteers: 228 scans; more than 40,000 ultrasound frames; straight line, C shape and S shape; distal-to-proximal direction; 36 ~ 430 frames for each scan; 100 ~ 200 mm per scan. | NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. The data was split into train, validation and test sets by a ratio of 3:1:1. | 1) Frame prediction accuracy: Euclidean distance between the predicted and ground truth positions of four corner points in consecutive frames; 2) Accumulated tracking error: average Euclidean distance over all reconstructed image pixel locations; 3) Volume reconstruction overlap: Dice coefficient between the reconstructed volumes of the prediction and the ground truth; 4) Final drift: Euclidean distance, averaged over the four corners, between the final predicted frame and the ground truth frame in each scan. | For a selected transformation (T_{10<--5}), without accumulated and consistency loss: CNN: 1) Frame prediction accuracy: 0.53±0.56 mm; 2) Accumulated tracking error: 19.09±14.60 mm; 3) Volume reconstruction overlap: 0.81±0.11; 4) Final drift: 36.82±28.01 mm; LSTM: 1) Frame prediction accuracy: 0.50±0.41 mm; 2) Accumulated tracking error: 18.13±9.49 mm; 3) Volume reconstruction overlap: 0.78±0.12; 4) Final drift: 36.48±20.77 mm
Privileged anatomical and protocol discrimination in trackerless 3D ultrasound reconstruction | Simplifying Medical Ultrasound | October 01, 2023 | Li et al. introduced a new multi-task learning framework for freehand ultrasound reconstruction, which leveraged anatomical and protocol information as privileged information and optimised the branching location of these auxiliary tasks via a differentiable architecture. | Sequences of M frames (M=100 or 140), sampled from 137 scans for training and 46 scans for validation. | A sequence of M frames (M=100 or 140) | No | EfficientNet (b1) | A number of selected transformations from C_M^2 transformations; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | 1) MSE between four corner points, transformed from predictions and ground truth, respectively. 2) Two cross-entropy losses between predicted class probability vectors and ground truth targets, for classifying anatomical and protocol. | No | Forearms from 19 volunteers: 228 scans; more than 40,000 US frames; straight line, C shape and S shape; distal-to-proximal direction; 36 ~ 430 frames for each scan; 100 ~ 200 mm per scan. | NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. The data was split into train, validation and test sets by a ratio of 3:1:1. | 1) Frame prediction accuracy: the Euclidean distance between the ground truth and prediction- transformed four corner points of each frame; 2) Accumulated tracking error: the averaged point distance on all pixels; 3) Volume reconstruction overlap: the overlap of all pixels between the ground truth and prediction volume; 4) Final drift: the frame prediction accuracy of the last frame in a scan. | M=100: 1) Frame prediction accuracy: 0.17±0.08 mm; 2) Accumulated tracking error: 3.80±3.97 mm; 3) Volume reconstruction overlap: 0.76±0.24; 4) Final drift: 6.56±7.53 mm; M=140: 1) Frame prediction accuracy: 0.13±0.05 mm; 2) Accumulated tracking error: 2.90±2.10 mm; 3) Volume reconstruction overlap: 0.89±0.00 4) Final drift: 6.53±5.98 mm;
Multi-IMU with online self-consistency for freehand 3D ultrasound reconstruction | Medical Image Computing and Computer Assisted Intervention | October 01, 2023 | Luo et al. proposed an online self-consistency network with multiple IMUs for improving reconstruction performance, along with a modal-level self-supervised strategy. | 1) N-length scanning sequences, sampled from 200 scans for training and 40 scans for validation (arm dataset), or from 150 scans for training and 30 scans for validation (carotid dataset); 2) M independent IMU data. | 1) N-length scanning sequence; 2) M independent IMU data. | No | ResNet + LSTM | M*(N-1) transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error (MAE) between prediction and ground truth; 2) Pearson correlation loss between prediction and ground truth; The prediction is obtained by averaging the M estimated parameters from repeatedly inputting the ultrasound images and M different IMU data into the backbone. Test phase: 1) Pearson correlation loss for measuring the difference between the estimated acceleration (obtained from the estimated transformation parameters) and IMU acceleration; 2) MAE for measuring the difference between the estimated angle (obtained from the estimated transformation parameters) and IMU angle; | Test phase: 1) Multi-IMU consistency constraints, including i) Pearson correlation loss for measuring the difference between the estimated accelerations; and ii) MAE loss for measuring the difference between the estimated angles. 2) Self-consistency constraint: MAE between the estimated 6-DoF from a sampled sub-sequence and the sampled 6-DoF from the original estimation. | 1) Arm dataset: 288 scans; linear, curved, loop, and sector scans; average length of 323.96 mm; 2) Carotid dataset: 216 scans; linear, loop, and sector scans; average length of 203.25 mm. | EM positioning system | Yes. 1) Arm dataset: training: 200 scans; validation: 40 scans; test: 48 scans; 2) Carotid dataset: training: 150 scans; validation: 30 scans; test: 36 scans. Randomly augmented (sub-sequence intercepting, interval sampling, and sequence inversion) each training scan to 20 sequences and each test scan to 10 sequences. | 1) Final drift rate (FDR); 2) Average drift rate (ADR); 3) Maximum drift (MD); 4) Sum of drift (SD); 5) Symmetric Hausdorff distance (HD); 6) Mean error of angle (EA) | 1) Arm dataset: FDR =10.01±5.7 % ; ADR = 18.86±6.5 %; MD = 28.61±11.0 mm; SD = 1064.06±582.5 mm; HD = 27.38±11.4 mm; EA = 2.76±1.3 deg; 2) Carotid dataset: FDR = 10.90±5.3 %; ADR = 19.61±8.5 %; MD = 21.81±7.2 mm; SD = 804.27±282.8 mm; HD = 19.30±7.6 mm; EA = 2.60±1.6 deg.
Trackerless volume reconstruction from intraoperative ultrasound images | Medical Image Computing and Computer Assisted Intervention | October 01, 2023 | Based on a Siamese architecture, EI hadramy et al. leveraged RNN and the optical flow for pose estimation. | Sequences of two images with 3 channels each, 2 heatmaps and an ultrasound frame. Sampled from 9614 clips for training and 2060 clips for validation. | Two frames (the first, middle frames or the middle, last frames of the ultrasound sequence with 7 frames) and 4 Gaussian heatmaps (obtained from sparse optical flow for 2 selected points), for each of the two Siamese networks. | Sparse optical flow for 2 selected points | Sequence to Vector (Seq2Vec) network | Two transformations (one is from the first frame to the fourth frame, and the other is from the fourth frame to the last frame); each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Mean square error (MSE) between the estimated transformations (from the first frame to the fourth frame, and from the fourth frame to the last frame) at each corner point of the frames and their respective ground truth; 2) MSE between the estimated transformation (from the first frame to the last frame, obtained by composing the two transformations above) at the corner points of the frames and the ground truth. | No | Ex vivo swine liver; six scans | Electromagnetic tracking system (trakSTAR™, NDI, Canada) | Yes. The data with 13734 clips was split into train, validation and test sets by a ratio of 7:1.5:1.5. | 1) Final drift rate (FDR): the final drift divided by the sequence length; 2) Average drift rate (ADR): the average cumulative drift of all frames divided by the length from the frame to the starting point of the sequence. | 1) FDR: 23.11±11.6 %; 2) ADR: 28.71±12.97 %; 3) Translation error: 0.449±0.189 mm; 4) Orientation error: 1.3±1.5 deg.
Long-term dependency for 3D reconstruction of freehand ultrasound without external tracker | IEEE Transactions on Biomedical Engineering | October 19, 2023 | Li et al. explored the impact of long-term dependencies on reconstruction performance, demonstrating how sequence length, anatomical content, and scanning protocol influence reconstruction quality, providing insights for optimising training data collection, scanning procedures, and network design. | Sequences of M frames, sampled from 137 scans for training and 46 scans for validation. (M=10,20,30,40,49,60,75,100) | A sequence of M frames for EfficientNet (b1) or one frame for LSTM | No | EfficientNet (b1) or EfficientNet (b1) + LSTM | τ+1 selected transformations from C_M^2 transformations; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation); τ+1 = 45,80,124,157,165,177,197,218 | MSE between four corner points, transformed from predictions and ground truth, respectively. | No | Forearms from 19 volunteers: 228 scans; straight line, C shape and S shape; distal-to-proximal direction; 36 ~ 430 frames each scan; 100 ~ 200 mm per scan. | Optical tracker (NDI Polaris Vicra, Northern Digital Inc., Canada) | Yes. The data was split into train, validation and test sets by a ratio of 3:1:1. This paper also tested the influence of anatomy and protocol to reconstruction performance, by training on one type of scans (e.g., straight line shape) and test on the original test set. | 1) Frame prediction accuracy: Euclidean distance between predicted and ground truth four corner points on the jth frame, both transformed from the ith frame; 2) Accumulated tracking error: average Euclidean distance of all reconstructed image pixels between prediction and ground truth; 3) Volume reconstruction overlap: the overlap between the ground-truth- and prediction- reconstructed scan volumes; 4) Final drift: the average Euclidean distance, over the four corner points on the last frame of the scan, between ground truth and prediction. | An example when M=100: 1) Frame prediction accuracy: 0.19±0.08 mm 2) Accumulated tracking error: 4.01±4.01 mm 3) Volume reconstruction overlap: 0.77±0.17 4) Final drift: 7.24±8.33 mm
Sensorless end-to-end freehand ultrasound with physics inspired network | IEEE International Ultrasonics Symposium | November 07, 2023 | To better leverage the contextual cues between frames, Dou et al. utilised speckle decorrelation as a physics-based constraint. | Two adjacent ultrasound images, sampled from 1,928 scans. | Two adjacent ultrasound images | No | ResNet | Relative transformation matrix, including translation, tilt, and rotation. | Mean square error (MSE) loss | No | Prostate dataset: 3D transrectal ultrasound (1,928 scans); | The freehand scans were resampled by generating oblique slices from the existing scans. | Yes | 1) Distance error; 2) Final drift | 1) Distance error: 4.38 mm; 2) Final drift: 7.06 mm.
IMU-assisted manual 3D-ultrasound imaging using motion-constrained swept-fan scans | Ultrasonic Imaging | May 2024 | Harindranath et al. presented an affordable IMU-assisted manual 3D ultrasound scanner, combining a consumer-grade IMU with Kalman filter-based orientation estimation and a scanline-based reconstruction method. | Data from IMU (only orientation) | NA | NA | NA | NA | NA | NA | 1) A water-filled table tennis ball of 40 mm diameter; 2) A commercial fetal phantom (SPACEFAN-ST, from Kyoto Kagaku, Japan); 106 swept-fan scan; 3) Data from two pregnant subjects. | EMT (Fastrak, Polhemus Ltd, Vermont, USA) | NA | 1) RMS error in angle 2) Diameter in axial, lateral, and elevational directions for data 1) 3) Biparietal diameter (BPD) and Gestational age for data 2) and 3) | 1) RMS error of 0.650° 2) For data 1): Axial: 40.4 ± 0.3 mm; Lateral: 40.4 ± 0.4 mm; Elevational: 40.2 ± 0.7 mm (manufacturer’s specifications of 40 mm (±0.004 mm)) 3) For data 2): BPD: absolute error of <2 mm; gestational age <1 week difference; 4) For data 3): absolute error of <2 mm in BPD and <1 week in gestational age measurements.
Sensorless volumetric reconstruction of fetal brain freehand ultrasound scans with deep implicit representation | Medical Image Analysis | May 2024 | Yeung et al. reconstructed 3D volumes from freehand 2D ultrasound sweeps without sensors by implicitly modeling the volume as a continuous function, while jointly refining slice locations and learning the full reconstruction. | 1) N 2D ultrasound images, sampled from a set of co-aligned 3D brain volumes; (for datasets A and B: non-uniformly sampled around the central axis of the brain; for native 2D freehand video sequences: randomly sampled from the sequences; N=128 or 256); 2) 3D locations, computed from estimated 3 Euler angles and 3 translations. | 1) N 2D ultrasound images (N=128 or 256); 2) 3D locations, computed from estimated 3 Euler angles and 3 translations. | Positional encoding | MLP | Intensity values of the corresponding pixels | Structural similarity (SSIM) loss between the predicted and actual intensity values | No | 1) Dataset A: fifteen 3D ultrasound fetal brain volumes; obtained as part of the INTERGROWTH-21st study; between 19 and 21 gestational weeks; 2) Dataset B: seventeen 3D ultrasound fetal brain volumes; between 19 and 21 gestational weeks; 3) Native 2D freehand video sequences: four videos of native freehand 2D brain scans; around 250 frames each; 20 weeks’ gestational age. | 1) For datasets A and B: ground truth is the original volumes; 2) For native 2D freehand video sequences: lack of ground truth 3D location information. | Yes. For native 2D freehand video sequences: 80% of each 2D ultrasound sequence for volumetric reconstruction and the remaining 20% for testing. | 1) Structural similarity index measure (SSIM); 2) Visual information fidelity (VIF): measuring the similarity between pair of 2D slices sampled from the 3D reconstructed volumes and the corresponding slices sampled from the native volumes; 3) Absolute difference between rotation angles (theta_diff) from refined locations and the ground truth locations; 4) Absolute distance between translations (T_diff) from refined locations and the ground truth locations; 5) Dice similarity coefficient (DSC); 6) 95th percentile Hausdorff distance (H_95); 7) Unsigned relative volume differences. | Boxplot; exact numbers are not provided.
Ultrasound volume reconstruction from 2D freehand acquisitions using neural implicit representations | IEEE International Symposium on Biomedical Imaging | August 22, 2024 | Gaits et al. proposed a 3D ultrasound reconstruction method based on neural implicit representations. | 3D spatial coordinates | 3D spatial coordinates | Positional encoding | NeRF (eight layers of 256 fully-connected perceptrons with ReLU activation) | Ultrasound slice pixel values | Mean squared error between the network output and ultrasound slice pixel value | No | 1) Simulated dataset: a synthetic medium consisting of a highly echogenic 1.25^3 cm^3 cube in a low echogenic homogenous medium; a sectorial scan (probe tilt); 2) Simulated dataset: a pelvic T2 MRI volume; double orthogonal scan; 3) A phantom mimicking a peripheral nerve block (CAE blue phantom from CAE Healthcare) | Magnetic tracker | Yes | 1) Signal-to-noise ratio (SNR); 2) Contrast-to-noise ratio (CNR) | 1) Cube: SNR: 31.8 dB; CNR: 30.4 dB; 2) MRI simulation: SNR: 28.3 dB; CNR: 25.8 dB; 3) Nerve phantom: SNR: 14.4 dB; CNR: 14.2 dB.
Sensorless end-to-end freehand 3-D ultrasound reconstruction with physics-guided deep learning | IEEE Transactions on Ultrasonics, Ferroelectrics, and Frequency Control | September 20, 2024 | Dou et al. utilised a two-stream model to separately extract spatial and temporal information, in addition to the incorporation of speckle decorrelation. | Sequences of four ultrasound frames, sampled from 45 cases for training and 15 cases for validation (phantom dataset), or from 137 cases for training and 45 cases for validation (forearm dataset). | A sequence of four ultrasound frames (randomly selected within the range of ±10 frames) | No | ResNeXt | Quaternion between input frames | 1) Mean squared error (MSE) between the estimated and ground truth transformation parameters; 2) Pearson correlation loss. | Triplet ranking loss, learning a discriminative feature embedding by penalising estimated samples based on their distance in feature space. | 1) Abdominal phantom: 75 scans; linear along the phantom contour; 30 ~ 94 mm; 2) Forearms from 19 volunteers: 228 scans; straight line, C shape and S shape; distal-to-proximal direction; 36 ~ 430 frames each scan; 100 ~ 200 mm per scan. | For data 1): A 6-D EM tracking system (NDI 3D Guidance, Northern Digital Inc., Waterloo, ON, Canada); For data 2): NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. For both datasets, the data were split into training, validation, and test sets using a 3:1:1 ratio. | 1) Distance error: average distance between predicted and ground truth frames over corner points in the ultrasound sequence; 2) Final drift: Euclidean distance of the final frame between the prediction and ground truth; 3) Frame error, providing the mean individual FE using the relative pose between prediction and ground truth; 4) MGU: GPU memory efficiency; 5) Average Inference Time (AIT). | Phantom dataset: 1) Distance error: 3.09±1.40 mm 2) Final drift: 4.65±2.25 mm 3) Frame error: 0.15±0.04 mm 4) MGU: 6050 MiB 5) Average Inference Time (AIT): 0.98 s Forearm dataset: 1) Distance error:12.77±6.84 mm 2) Final drift: 18.01±10.44 mm 3) Frame error: 0.48±0.26 mm 4) MGU: 6932 MiB
Geometric transformation uncertainty for improving 3D fetal brain pose prediction from freehand 2D ultrasound videos | Medical Image Computing and Computer Assisted Intervention | October 03, 2024 | Ramesh et al. proposed an uncertainty-aware deep learning model for 3D plane localisation in 2D fetal brain images, using a multi-head network that regressed different geometric transformations and uncertainty. | N 2D ultrasound images, sampled from 4 training and 2 validation 3D volumes aligned to the 3D anatomical atlas. | N 2D ultrasound images | No | 2D CNN | Five sets of poses and variances for each input image: quaternions, axis-angles, euler angles, rotation matrices, translation displacement and scaling, and direct pose (3 Cartesian coordinates of 3 reference locations (the center, bottom right, and bottom left of the plane)) | 1) Mean-squared-error (MSE) loss 2) Heteroscedastic Gaussian Negative Log-Likelihood | No | 3D ultrasound fetal brain volumes: obtained as part of the INTERGROWTH-21st study; 4 training, 2 validation, and 7 test 3D volumes selected at 19 gestational weeks. | Pre-aligned to a common reference atlas space; Artificial 2D slice sampling (training labels derived from the nature of the data itself) | Yes | 1) Euclidean distance (ED) and plane angle (PA) between all coordinates of the predicted and ground truth planes measured in radians; 2) Mean squared error (MSE) between predicted and ground truth reference points; 3) Normalized Cross-Correlation (NCC) and Structural Similarity (SSIM) between images sampled with predicted planes and ground truth planes. | 1) ED: 0.32±0.24 2) PA: 0.39±0.24 3) MSE: 215.10±241.52 4) NCC: 0.67±0.22 5) SSIM: 0.61±0.25
Fine-grained context and multi-modal alignment for freehand 3D ultrasound reconstruction | Medical Image Computing and Computer Assisted Intervention | October 03, 2024 | Yan et al. proposed ReMamba, featuring a multi-directional state space model for multi-scale spatio-temporal mining, an adaptive fusion strategy using multiple IMUs for enhanced perception, and an online alignment strategy leveraging pseudo labels for improved reconstruction. | 1) N-length scanning sequence, sampled from 375 cases for training and 104 cases for validation (arm dataset), or from 276 cases for training and 78 cases for validation (carotid dataset); 2) M independent IMU data. | 1) N-length scanning sequence; 2) M independent IMU data | No | ReMamba, consisting of stem, multiple residual blocks and ReMamba blocks with multi-directional SSM. | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error between network prediction and ground truth; 2) Pearson correlation loss between network prediction and ground truth; Test phase: 1) Alignment loss (mutual information between image feature and weighted temporal feature of IMU acceleration); 2) Pearson correlation loss between the estimated Euler angle and weighted IMU angle. | No | 1) Arm dataset: 583 scans; linear, curved, loop, and sector scans; average length of 386.39 mm; 2) Carotid dataset: 432 scans; linear, loop, and sector scans; average length of 241.25 mm. | EM positioning system | Yes. Arm dataset: training: 375 scans; validation: 104 scans; test: 104 scans; Carotid dataset: training: 276 scans; validation: 78 scans; test: 78 scans. Performed random augmentations on each scan (sub-sequence intercepting, interval sampling, and sequence inversion). | 1) Final drift rate (FDR); 2) Average drift rate (ADR); 3) Maximum drift (MD); 4) Sum of drift (SD); 5) Symmetric Hausdorff distance (HD); 6) Mean error of angle (MEA) | 1) Arm dataset: FDR = 9.72±7.1 %; ADR = 15.53±9.6 %; MD = 24.68±13.6 mm; SD = 677.48±498.4 mm; HD = 23.05±13.6 mm; MEA = 3.41±1.8 deg; 2) Carotid dataset: FDR = 8.61±5.9 %; ADR = 16.16±8.1 %; MD = 13.78±6.4 mm; SD = 391.77±298.4 mm; HD = 12.39±6.2 mm; MEA = 2.09±1.2 deg.
Advancing sensorless freehand 3D ultrasound reconstruction with a novel coupling pad | Medical Image Computing and Computer Assisted Intervention | October 14, 2024 | Dai et al. developed a novel coupling pad with 3 N-shaped lines to provide 3D spatial information without external tracking devices and introduced a coarse-to-fine optimisation method that refined sequential 2D ultrasound image poses via distance-topology discrepancy reduction. | A sequence of ultrasound images | NA | NA | NA | NA | 1) Topology discrepancy between the simulated points (generated by the N-shaped lines passing through the image plane) and the corresponding detected points in the image; 2) Distance discrepancy between the simulated points and the corresponding detected points in the image. | No | 1) Synthetic sequential 2D ultrasound images with markers: generated from 3D images; 100 linear scans; 80 ~ 100 frames per scan; scanning lengths of 65 ~ 80 mm; 2) 9-line phantom: 80 ~ 150 frames per scan; 45 scans; linear, oscillating, and back-and-forth; 3) Solid phantom (melamine sponge): 80 ~ 150 frames per scan; 15 scans; linear, oscillating, and back-and-forth. | For dataset 1): simulation with known ground truth; For dataset 2): theoretical lines; For dataset 3): known geometry relationships | NA | For data 1): Final Drift Rate (FDR), Average Drift Rate (ADR), Maximum Drift (MD), Sum of Drift (SD), and Symmetric Hausdorff Distance (HD) For data 2): mean distance from reconstructed 9 lines and theoretical lines For data 3): mean reconstruction errors | 1) FDR: 2.74±2.98 %; ADR: 3.35±3.24 %; MD: 2.52±2.25 mm; SD: 100.29±103.02 mm; HD: 2.05±2.06 mm 2) < 0.5 mm 3) Linear scans: 0.33±0.20 mm; oscillating scans: 0.52±0.31 mm; back-and-forth scans: 0.28±0.23 mm.
Nonrigid reconstruction of freehand ultrasound without a tracker | Medical Image Computing and Computer Assisted Intervention | October 14, 2024 | Li et al. proposed a co‑optimisation framework that jointly estimated rigid and non‑rigid transformations, with options for meta‑learning or weighted objectives, and a fast scattered‑data interpolation. | Sequences of 100 ultrasound frames, sampled from 432 scans for training and 144 scans for validation. | A sequence of 100 ultrasound frames | No | EfficientNet (b1) + Voxelmorph | From EfficientNet (b1): 99 transformations between each frame and the first frame; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation); From Voxelmorph: dense displacement field which transformed the rigid-reconstruction-predicted volume to the final prediction. | 1) Mean squared error between predicted points coordinates and ground truth points coordinates; 2) Intensity similarity between the ground-truth-reconstructed volume and wrapped predicted volume. | Bending energy of predicted deformation. | Forearm dataset from 60 volunteers: 720 scans; straight, c-shape and s-shape; distal-to-proximal direction; average number of frames per scan (500); 200~300 mm per scan. | NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. The dataset was split into train, validation and test sets by a ratio of 3:1:1. | 1) Global pixel reconstruction error (GPE), reconstruction error on all pixels based on global displacement vectors; 2) Global landmark reconstruction error (GLE), reconstruction error on landmarks based on global displacement vectors; 3) Local pixel reconstruction error (LPE), reconstruction error on all pixels based on local displacement vectors; 4) Local landmark reconstruction error (LLE), reconstruction error on landmarks based on local displacement vectors. | 1) GPE: 16.51±7.76 mm; 2) GLE: 17.91±7.85 mm; 3) LPE: 3.07±0.99 mm; 4) LLE: 3.38±1.00 mm.
Pitfalls with neural radiance fields for 3D freehand ultrasound reconstruction | Annual International Conference of the IEEE Engineering in Medicine and Biology Society (EMBC) | December 17, 2024 | Dou et al. addressed pixel mismatches in freehand ultrasound NeRF caused by transducer pressure variations, introducing a modified positional encoding that improved 3D representation learning. | 3D positions of ultrasound images (randomly sampled 16384 rays at each learning step) from four scans of a volume | 3D positions of query points | Positional encoding | Ultra-NeRF | A parameter vector (attenuation, reflectance, boundary probability, scattering density, scattering intensity), which can be rendered into gray scale intensity | L2 loss between the generated images with their corresponding target images | No | Two freehand datasets on a tissue mimicking abdominal phantom (Model 057A, CIRS Inc., USA) 1) Four scans; 621 frames; 2) Four scans; 865 frames; | A 6D electromagnetic (EM) tracking system (NDI 3D Guidance, Northern Digital Inc., Canada) | Yes | 1) Peak signal to noise ratio (PSNR); 2) Structural similarity index (SSIM) | 1) Phantom Data 1: PSNR: 27.17 dB; SSIM: 0.56 %; 2) Phantom Data 2: PSNR: 26.11 dB; SSIM: 0.51 %.
Continuous Bernoulli distribution for more realistic ultrasound reconstruction with NeRF | IEEE Ultrasonics, Ferroelectrics, and Frequency Control Joint Symposium | December 18, 2024 | Dou et al. proposed a probabilistic and differentiable method that used a continuous Bernoulli distribution to model scatterer distributions and backscattered pixel intensities, enabling more realistic 3D ultrasound reconstruction with NeRF. | 3D positions from 800 two-dimensional B-mode liver images (randomly sampled 2000 rays at each learning step) | 3D positions of query points, randomly sampled 2000 rays at each learning step | No | Ultra-NeRF | A parameter vector (attenuation, reflectance, boundary probability, scattering density, scattering intensity), which can be rendered into gray scale intensity | L2 loss between the generated images with their corresponding target images | No | Synthetic liver data: 800 two-dimensional B-mode liver images for training. | Not clear | Yes | 1) Peak signal-to-noise ratio (PSNR); 2) Structural similarity index (SSIM) | 1) PSNR: 27.73±0.98 dB 2) SSIM: 0.54±0.19
3d freehand ultrasound reconstruction by reference-based point cloud registration | International Journal of Computer Assisted Radiology and Surgery | March, 2025 | Großbröhmer et al. proposed a three-stage pipeline for trajectory estimation in freehand ultrasound, combining a Free Point Transformer, HeatReg refinement with support point clouds, and instance optimisation. | 1) Two ultrasound scans (fixed and moving), sampled from 19 training, 1 validation case (20 reference, 40 moving), using a leave-one-out cross-validation; 2) Rotational IMU data (initialise the registration using rotational IMU data) | Two point clouds (reference and moving) | 1) Automatically generated two bone surface point clouds (2048 points each); 2) Support keypoints; 3) Gradient keypoints | Free Point Transformer (FPT) + HeatReg + instance optimisation | Displacement field (2048*3) | For registration: 1) difference between warped moving from FPT and tracked ground truth; 2) difference between final prediction from HeatReg and tracked ground truth; For instance optimisation: 1) Chamfer Loss between predicted and fixed point cloud; 2) DiVRoC loss. | For registration: L1 Loss with equal weighting on predicted displacements both with and without translational regularization. | 63 left forearm scans from 21 subjects: lateral direction; straight dorsal, C-shape dorsal and S-shape dorsal; length of a scan varies between 67 and 279 frames (157±42); a length of 11.6-17.3 cm (14.7±1.6) | NDI Hybrid Polaris Spectra (Northern Digital Inc, Canada) | Yes. “straight dorsal” as fixed and “C-shape”/“S-shape” as moving sweeps. Leave-one-out cross-validation with 19 training, 1 validation and 1 testing case | 1) Directed Surface Distance (DSD): the average distance between each point in the transformed point cloud and its nearest neighbour in the fixed point cloud; 2) Target registration error (TRE): assessment of direct point correspondences, computed with the ground truth trajectories of moving sweeps; 3) Average symmetric surface distance (ASSD), computed with the ground truth trajectories of moving sweeps; 4) 95th percentile of Hausdorff Distances (HD95), computed with the ground truth trajectories of moving sweeps; 5) Translational pose error (E_T): sweep trajectory recovery, computed with the ground truth trajectories of moving sweeps. 6) Runtime | 1) DSD: 1.29±0.38 mm 2) TRE: 5.20±2.15 mm 3) ASSD: 1.13±0.21 mm 4) HD95:2.59±0.80 mm 5) E_T: 6.14±2.66 mm 6) Runtime: 880 ms
Rapidvol: rapid reconstruction of 3D ultrasound volumes from sensorless 2D scans | IEEE International Symposium on Biomedical Imaging | May 12, 2025 | Eid et al. proposed a neural representation framework for fast slice-to-volume ultrasound reconstruction using tensor-rank decomposition. | 3D coordinates of all pixels on a set of 2D ultrasound images (128 and 256 axial or coronal slices linearly sampled from the 3D scan) | 3D coordinates of all pixels on one 2D ultrasound image, obtained using the pose of the image. | Positional encoding | Tri-Planar Decomposition + MLP | A grayscale image at the requested pose | Negative structural similarity index measure between the rendered image and the ground truth image | No | 15 3D US fetal brain scans: 20 gestational weeks; collected from the INTERGROWTH-21st study | Ground truth poses are sampled from the 3D scan | Yes | Structural similarity index measure (SSIM) | SSIM results: 1) Axial input training scan sets and 128 testing slices: Coronal: .955 ± .008; Sagittal: .952 ± .009; 2) Axial input training scan sets and 256 testing slices: Coronal: .973 ± .005; Sagittal: .969 ± .006; 3) Coronal input training scan sets and 128 testing slices: Axial: .941 ± .016; Coronal: .932 ± .018; Sagittal: .935 ± .017; 4) Coronal input training scan sets and 256 testing slices: Axial: .952 ± .015; Coronal: .950 ± .016; Sagittal: .950 ± .014.
MoNetV2: Enhanced motion network for freehand 3-D ultrasound reconstruction | IEEE Transactions on Neural Networks and Learning Systems | June 11, 2025 | As an extension of Luo et al. (2022), Luo et al. proposed scan-level, path-level, and patch-level consistency to enhance the reconstruction accuracy. | 1) N-length scanning sequences, sampled from 7500 generated sequences from 375 scans in each training epoch and 1040 generated sequences from 104 scans for validation (arm dataset), or from 5520 generated sequences from 276 scans in each training epoch and 780 generated sequences from 78 scans for validation (carotid dataset); or from 3680 generated sequences from 184 scans in each training epoch and 530 generated sequences from 53 scans for validation (thyroid dataset); 2) IMU data | 1) All pairs of adjacent images from a N-length scanning sequence; 2) IMU acceleration; 3) IMU Euler angle. | No | ResNet18 + ConvLSTM | N-1 transformations between adjacent frames; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | Training phase: 1) Mean absolute error (MAE) between the estimated transformation parameters and the ground truth; 2) Pearson correlation loss between the estimated transformation parameters and the ground truth; Test phase: Online multimodal self-supervised strategy: Pearson correlation loss between the estimated acceleration (obtained from the estimated translation) and IMU acceleration; MAE loss between the estimated Euler angle and IMU Euler angle. | Test phase: 1) Scan-level velocity consistency: velocity constrains between predictions with different intervals (the MAE between long-interval transformations and the composition of the shorter-interval transformations); 2) Path-level appearance consistency: differences between transformations of reordered subsequences and the reordered original predictions; 3) Patch-level motion consistency: MAE between the normalised content difference and the normalised estimated distance, for two adjacent images. | 1) Arm dataset: 583 scans; average length of 386.39 mm; linear, curved, loop, and sector scans; 2) Carotid dataset: 432 scans; average length of 241.25 mm; linear, loop, and sector scans; 3) Thyroid dataset: 290 scans; average length of 66.25 mm; linear and loop scans. | EM positioning system (3-D Guidance trakSTAR, Northern Digital Inc., Canada) | Yes. Arm dataset: training: 375 scans; validation: 104 scans; test: 104 scans; Carotid dataset: training: 276 scans; validation: 78 scans; test: 78 scans; Thyroid dataset: training: 184 scans; validation: 53 scans; test: 53 scans. Augmented (subsequence intercepting, interval sampling, and sequence inversion) each training scan into 20 different sequences in each training epoch; Augmented each validation or test scan into 10 fixed sequences. | 1) Final drift rate (FDR): the ratio of the distance between the estimated and actual positions of the last frame divided by the scan length; 2) Average drift rate (ADR): the mean drift rate, where each drift rate is defined as the ratio of a frame’s cumulative drift to its distance from the initial frame of the scan; 3) Maximum drift (MD): the MD among all the frames; 4) Sum of drift (SD): the total sum of all the drifts; 5) Symmetric Hausdorff distance (HD): the largest distance between estimated and actual positions across all frames; 6) Mean error of angle (MEA): the MAE between the estimated Euler angle and the actual Euler angle. | 1) Arm dataset: FDR: 11.04±6.02 %; ADR: 18.26±9.10 %; MD: 28.84±15.62 mm; SD: 901.87±658.43 mm; HD: 26.24±12.47 mm; MEA: 3.00±2.25 deg; 2) Carotid dataset: FDR: 12.11±7.55 %; ADR: 23.78±17.51 %; MD: 20.75±8.82 mm; SD: 646.00±455.03 mm; HD: 18.50±8.64 mm; MEA: 2.40±1.55 deg; 3) Thyroid dataset: FDR: 17.88±9.10 %; ADR: 26.34±9.12 %; MD: 8.93±4.79 mm; SD: 207.60±195.20 mm; HD: 8.43±4.73 mm; MEA: 1.36±0.97 deg.
Enhancing free-hand 3D photoacoustic and ultrasound reconstruction using deep learning | IEEE Transactions on Medical Imaging | June 13, 2025 | Lee et al. proposed a motion-based learning network with a global-local self-attention module, incorporating correlation features, global-local attention, and a motion-aware supervision strategy to enhance motion estimation. | Two ultrasound sequences with s+1 frames, sampled from 30 scans for training and 6 scans for validation (for in-house forearm dataset), or from 30 scans for training and 6 scans for validation (for in-house Photoacoustic dataset), or from 120 scans for training and 48 scans for validation (for public forearm dataset). | Two ultrasound sequences with s+1 frames (the first sequence includes frames n to n+s; the second sequence includes frames n+1 to n+s+1) | No | CNN + RNN + self-attention mechanism | s+1 transformations; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) | 1) Motion-based mean absolute error: assigns higher weight to fast motion; 2) Correlation loss between predicted 6-DoFs and ground truth. | Margin triplet loss: encourages the model to contrast the feature maps in latent space. | 1) Forearms from 9 individuals: 54 scans; 900~1000 frames per scan; 51,300 frames in total; an arbitrary S-shaped trajectory; three scans for each forearm; travel length of 199.27±23.81 mm; 2) Photoacoustic data: 60 scans; 10 volunteers; 28,445 frames in total; travel length of 153.12±13.20 mm; 3) Forearms from 19 volunteers: 228 scans; 30,552 frames in total; straight line, C shape and S shape; distal-to-proximal direction; 36 ~ 430 frames each scan; travel length of 144.51±14.26 mm. | For data 1) and 2): Electromagnetic sensor (Polhemus); For data 3): NDI Polaris Vicra (Northern Digital Inc., Canada) | Yes. For dataset 1): training: 30 scans; validation: 6 scans; test: 18 scans; For dataset 2): The dataset was split into training, validation, and testing in a 5:1:4 ratio. For dataset 3): The dataset was divided into training, validation, and testing in a 10:4:5 ratio. During inference, only the final prediction is used to reconstruct the 3D volume. | 1) Relative average error (rAE): measures the relative motion estimation accuracy; 2) Accumulated average error (aAE): measures the accumulated motion error over the scan trajectory; 3) Relative frame error (rFE): relative displacements between frames; 4) Accumulated frame error (aFE): cumulative displacements along the scan path; 5) Correlation (Corr): cosine similarity, to measure the underlying trends of reconstructed trajectories in 3D Euclidean space; 6) Final drift rate (FDR): aFE of the final frame divided by the total length of the scan trajectory. | 1) Dataset 1: rAE: 0.1430±0.072 mm, 0.0663±0.023 deg; aAE: 7.9097±4.89 mm, 5.9412±3.83 deg; rFE: 0.2917±0.054 mm; aFE: 16.1944±6.09 mm; Corr: 0.9217±0.048; FDR: 12.0577±6.45 %. 2) Dataset 2: rAE: 0.0597±0.026 mm, 0.0221±0.008 deg; aAE: 4.3204±3.31 mm, 2.2230±1.88 deg; rFE: 0.1225±0.042 mm; aFE: 8.9360±3.99 mm; Corr: 0.9165±0.059; FDR: 10.6663±3.87 %; 3) Dataset 3: rAE: 0.1979±0.124 mm, 0.0755±0.026 deg; aAE: 8.0867±6.12 mm, 2.3956±1.57 deg; rFE: 0.4113±0.159 mm; aFE: 16.4062±7.33 mm; Corr: 0.9164±0.105; FDR: 19.8096±10.30 %.
UltrasOM: A mamba-based network for 3D freehand ultrasound reconstruction using optical flow | Computer Methods and Programs in Biomedicine | August 2025 | Sun et al. developed a Mamba-based spatio-temporal attention module, integrated with optical flow, to capture global spatio-temporal correlations. | Sequences of 8 frames, sampled from 160 training ultrasound scans and 20 validation ultrasound scans during training. | 1) A sequence of 8 frames; 2) Optical flow. | 1) 7 optical flow features; 2) spatial position encoding; 3) temporal position encoding. | Space-Time Blocks (Mamba blocks and space attention mechanisms) + Regression head | 7 transformations; each transformation has 6 degree of freedoms (3 for translation, 3 for rotation) and represents the relative pose between each adjacent frame pair | 1) Mean squared error (MSE): the average difference between predicted and ground truth 6-DoF transformation parameters across the n-1 relative transformations; 2) Correlation loss based on Pearson's correlation coefficient; 3) Motion speed loss: the average speed error between predicted and ground truth velocities for consecutive frames, computed from translation and rotation information and averaged over the n−2 speed values. | No | 200 ultrasound video sequences of forearms from 20 volunteers: 58,011 frames; linear, curved, fast-slow, roll; an average scan length of 103 mm. | An optical tracking device, the NDI Polaris Vega XT (Northern Digital Inc., Canada) | Yes. The dataset was split into training, validation, and testing sets in an 8:1:1 ratio. Using sliding window reconstruction strategy during inference, which takes the average of these repeated predictions to obtain the final result for each frame. | 1) Distance Error (DE): the average distance between corresponding corner points across all frames in the sequence; 2) Final Drift (FD): the positional difference between the last frame of the reconstructed sequence and the ground truth; 3) Final Drift Rate (FDR): defined as FD divided by the sequence length, where the sequence length is calculated by connecting the ground truth center points of each frame; 4) Average Drift Rate (ADR): the average drift per frame; 5) Maximum Drift (MD): the largest drift within the sequence; 6) Sum of Drift (SD): total drift across all frames; 7) Symmetric Hausdorff Distance (HD): the discrepancy between the true and estimated displacement parameters; 8) Mean Angular Error (MEA): the average absolute error between the true and estimated rotational parameters. | 1) DE: 7.34±3.11 mm; 2) FD: 10.67±5.91 mm; 3) FDR: 10.24±5.34 %; 4) ADR: 17.15±4.60 %; 5) MD: 13.15±5.50 mm; 6) SD: 1966.59±784.56 mm; 7) HD: 10.81±5.81 mm; 8) MEA: 2.05±1.74 deg.